YSI Model 85 Dissolved Oxygen Meter

The Model 85’s microprocessor allows the system to be easily calibrated for dissolved oxygen or conductivity with the press of a few buttons. Additionally, the microprocessor performs a self diagnostic routine each time the instrument is turned on. The self-diagnostic routine provides you with useful information about the conductivity cell constant and function of the instrument circuitry.

The system simultaneously displays temperature (in °C), along with one of the following parameters: dissolved oxygen in either mg/L (milligrams per liter) or % air saturation; conductivity; temperature compensated conductivity; (in μS/cm or mS/cm), and salinity (in parts per thousand {ppt}).

The system requires only a single calibration regardless of which dissolved oxygen display you use.  The calibration of conductivity is not required but is available. A single calibration will adjust the instrument, regardless if you are reading conductivity or temperature compensated conductivity.

YSI Model 85 Calibration of Dissolved Oxygen

To accurately calibrate the YSI Model 85 you will need to know the approximate altitude of the region in which you are located.  Calibrate the meter before you go to the field and leave it on until done for the day.  Do not let sponge inside the probe chamber dry out! Keep sponge moist by re‐wetting at least biweekly.  If the sensor membrane becomes damaged, you will need to replace it and wait 12 hours before calibrating and using the monitor.

1. Ensure that the sponge inside the instrument's calibration chamber is wet. Insert the probe into the calibration chamber.

2. Turn the instrument on by pressing the ON/OFF button on the front of the instrument.

Press the MODE button until dissolved oxygen is displayed in mg/L or %. Wait for the dissolved oxygen and temperature readings to stabilize (usually 15 minutes is required).

3. Record date, time, name of analyst calibrating the monitor, and the altitude (ft) where calibrating.

4. Record the number of minutes of warm up time.

5. Record the stabilized probe temperature (°C).

7. Press the MODE button to change from % saturation to (mg/L).

8. Record the stabilized pre‐calibration D.O. (mg/L).

9. Use two fingers to press and release both the UP ARROW and DOWN ARROW buttons at the same time.

10. The LCD will prompt you to enter the local altitude in hundreds of feet. Use the arrow keys to increase or decrease the altitude. When the proper altitude appears on the LCD, press the ENTER button once.  EXAMPLE: Entering the number 12 here indicates 1200 feet.

11. The Model 85 should now display CAL in the lower left of the display, the calibration value should be displayed in the lower right of the display and the current % reading (before calibration) should be on the main display. Make sure that the current % reading (large display) is stable, then press the ENTER button. The display should read SAVE then should return to the Normal Operation.

Press the MODE key to switch to (mg/L). (This key is used to toggle between mg/L and % saturation.)

12. Record the post‐calibration D.O. (mg/L).

13. Look up the calibration chart D.O. value (mg/L) from the table included with your meter.

Find the probe temperature and altitude and record the corresponding D.O. (mg/L) value as the calibration chart D.O. on the calibration log. If the difference between the post calibration

D.O. and the calibration chart D.O. is greater than 0.3 (mg/L), re‐calibrate the meter before using in the field.

Leave meter on until last reading of the day is completed.

14. Press the green button to turn the meter off at the end of the day.

Each time the Model 85 is turned off, it may be necessary to re-calibrate before taking measurements. All calibrations should be completed at a temperature which is as close as possible to the sample temperature. Dissolved oxygen readings are only as good as the calibration.

YSI Model 85 Calibration of Conductivity

IMPORTANT: System calibration is rarely required because of the factory calibration of the YSI Model 85. However, from time to time it is wise to check the system calibration and make adjustments when necessary.

Prior to calibration of the YSI Model 85, it is important to remember the following:

1. Always use clean, properly stored, NIST traceable calibration solutions (see Accessories and Replacement Parts). When filling a calibration container prior to performing the calibration procedures, make certain that the level of calibrant buffers is high enough in the container to cover the entire probe. Gently agitate the probe to remove any bubbles in the conductivity cell.

2. Rinse the probe with distilled water (and wipe dry) between changes of calibration solutions.

3. During calibration, allow the probe time to stabilize with regard to temperature (approximately

60 seconds) before proceeding with the calibration process. The readings after calibration are  only as good as the calibration itself.

4. Perform sensor calibration at a temperature as close to 25°C as possible. This will minimize any temperature compensation error.

Follow these steps to perform an accurate calibration of the YSI Model 85:

1. Turn the instrument on and allow it to complete its self-test procedure.

2. Select a calibration solution that is most similar to the sample you will be measuring.

• For sea water choose a 50 mS/cm conductivity standard (YSI Catalog# 3169)

• For fresh water choose a 1 mS/cm conductivity standard (YSI Catalog# 3167)

• For brackish water choose a 10 mS/cm conductivity standard (YSI Catalog # 3168)

3. Place at least 3 inches of solution in a clean glass beaker.

4. Use the MODE button to advance the instrument to display conductivity.

5. Insert the probe into the beaker deep enough so that the oval-shaped hole on the side of the probe is completely covered. Do not rest the probe on the bottom of the container -- suspend it above the bottom at least 1/4 inch.

6. Allow at least 60 seconds for the temperature reading to become stable.

7. Move the probe vigorously from side to side to dislodge any air bubbles from the electrodes.

8. Press and release the UP ARROW and DOWN ARROW buttons at the same time.

The CAL symbol will appear at the bottom left of the display to indicate that the instrument is now in Calibration mode.

9. Use the UP ARROW or DOWN ARROW button to adjust the reading on the display until it matches the value of the calibration solution you are using.

10. Once the display reads the exact value of the calibration solution being used (the instrument will make the appropriate compensation for temperature variation from 25°C), press the

ENTER button once. The word “SAVE” will flash across the display for a second indicating that the calibration has been accepted.

The YSI Model 85 is designed to retain its last conductivity calibration permanently. Therefore, there is no need to calibrate the instrument after battery changes or power down.

Oakton Acorn pH Meter Calibration Procedure

1. Connect pH and temperature probes to the meter and ensure the pH probe connecter snaps  into place.

2. Turn the meter on.

3. Use provided pH 7 and pH 10 buffer solutions to calibrate the meter.

4. Record the date, time, person doing calibration, and the pH 7 buffer standard value on the log sheet.

5. Remove electrode from container with storage solution.

6. Rinse electrode with distilled water and shake excess water off electrode.

7. Press CAL button on the meter to start the calibration process.

8. Place electrode and temperature sensor into the lower pH buffer solution and submerge the glass bulb completely.

9. Wait for reading to stabilize and record the Stabilized pH on the calibration log sheet.

10. Press the ENTER key once and record the Calibrated pH value.

11. Check that the Calibrated Value is within 0.2 of the pH buffer standard. If the value is more than 0.2 higher or lower than its pH buffer value, recalibrate the meter. Record new results on the next line of the Calibration Log.

12. Remove electrode from pH 7 buffer solution, rinse with distilled water and shake off excess water.

13. Repeat steps 7 through 12 for the pH 10 buffer solution.

14. Replace pH probe in storage solution, turn meter off, and head out to the field.

Note: If you notice your meter is not calibrating properly, try cleaning the probe by following instructions in the instrument’s manual (located in the pocket of your meter case) or contact the program coordinator for directions.

YSI Model 85 Dissolved Oxygen Meter Measurement

To Measure Dissolved Oxygen in the Stream:

1. Enter the stream, approaching the site you will monitor from downstream.

2. Insert the probe into the water to be measured.

3. Continuously stir or move the probe through the water (especially when monitoring in very still water).

4. Allow the temperature and dissolved oxygen readings to stabilize (about 2‐3 minutes).

5. Observe and record temperature and dissolved oxygen values on the monitoring data sheet.

6. Press MODE and observe and record the % saturation on the monitoring data sheet.

7. Rinse the probe with distilled water after each use, and replace in the probe chamber

8. Be sure to always keep the probe sponge moist by soaking with distilled water.

YSI Model 85 Conductivity Sampling Procedure

1.  While YSI Model 85 probe is in the stream, toggle from dissolved oxygen (% saturation to mg/L to mS.  Additionally the small portion of the display will show the o C NOT flashing.

Oakton Acorn pH Meter Sampling Procedure

1. Wade into stream mid‐width.

2. Face upstream and dip plastic sample bottle into stream with the mouth facing downward.

3. Submerge the bottle 12 inches or halfway between the surface and stream bottom.

4. Fill bottle by turning bottle (while underwater) so that its mouth faces up.

5. Rinse and dump water from sample bottle downstream.

6. Repeat steps 2‐4.

7. Transport sample bottle back to shore or vehicle.

8. Connect any electrodes to pH meter and turn meter on.

9. Remove electrode from container with storage solution.

10. Rinse electrode with distilled water and shake off excess water.

11. Place pH and temperature probes into sample bottle and gently stir the probes.

12. Wait 2‐3 minutes for pH reading to stabilize.

13. Record the pH reading on the log sheet. Do not record the temperature reading from the pH meter since it is not the actual stream temperature.

14. Rinse electrode with distilled water and shake off excess water.

15. Place electrode back in container with storage solution.

16. Disconnect electrodes and turn the meter off to store.

Note: Store pH probe in pH storage solution only. Never use distilled water. If you are out of solution, contact the program coordinator.

Transparency Monitoring Procedure

Sample Collection

Collect the sample away from the stream bank in the main flow (well‐mixed) area. Be careful not to disturb the stream bottom when you collect the water sample. If you get sediment from bottom disturbances, dump out the sample, move upstream away from the disturbed area and try again.

Never wear sunglasses when you take this reading.

In Stream

1.  Walk into the water downstream from the sampling location. Be careful not to stir up the bottom sediment upstream of your sampling location.

2.  Face upstream (into the current) in the middle of the stream or in a well‐mixed area offshore.

3. Collect your water sample by plunging your bucket or tube 8‐12 inches beneath the surface or halfway down from the surface. Scoop away from your body and into the current.

4.  Return to shore with the sample.

Reading the Transparency Tube

1. Remove large objects from the water sample. (Filter through nylon stocking if necessary.)

2. If the sample has settled, use a stirring stick to stir the sample, or pour the sample into a clean bucket and back into the transparency tube to suspend all materials.

3. Stand out of direct sunlight. If you cannot get to a shady place, use your body to cast a shadow on the tube (Figure 1).

4. If you are wearing sunglasses, remove them. Then look for the target (black and white) disc on the bottom of tube. If disc is visible, record the length of the tube (e.g., 120 cm) on the data sheet.

5. If target disc is not visible, have your partner let water out a little at a time using the valve at the bottom until disc is just visible (Figure 2). Have them stop letting water out immediately when you can just see the contrast between black and white on the disc.

6. Read the level of water in the tube in cm using the measuring tape on the side of the tube.

7. Record the measurement on your data sheet in cm.

8. Dump contents of tube on ground.

9. Collect a new sample then repeat steps 1 through 8.

10. Record the second measurement in cm on your data sheet.

Continuous Temperature Monitoring Procedure

Attempt to place the thermistor in an area that will be submerged, well‐mixed, and free of sedimentation during the period of deployment. If the thermistor gets buried by sediment or vegetation it will not be exposed to the true water temperature, and may not sense changes in temperature as quickly. A riffle or run is preferred to a pool. Seek shade so as to minimize any radiant heat from the sun. A deep run or riffle in the sun would be acceptable. Be sure to consider the safety of others using the stream and try your best to pound rebar stakes in where it is unlikely people will run into them.

Deploying your Thermistor in the Stream

When deploying your thermistor remember the two main goals are to 1) keep the thermistors submerged for the whole season and 2) to avoid the thermistor from getting buried by sediment or debris. A very important third goal though is to secure it well so that the thermistor is not lost (keep in mind the flow of your stream can increase drastically after heavy rain events).

1.  Connect the thermistor to railroad plate with zip ties. We recommend using two zip ties, or one zip tie and some wire, in case one of the zip ties fails. A zip tie around the rebar above the thermistor can also help the thermistor from sliding off the end of the rebar.

2.  Make sure the thermistor is a few inches above the stream bottom.

3.  Check to see if the light is blinking (you may need to shield the device to see it if the light is faint). If your device is not blinking it may not be logging data.

4.  Take an accurate GPS reading or make detailed notes about where you place the

thermistor in the stream (so you can find it again). Record this information, along with the device serial number, date, and time you placed the thermistor in the stream on your Thermistor Log.

During the Season

Check the thermistors each month and note the dates of your checks on the Thermistor Log. Check thermistor is not covered by sediment or vegetation and that it is still under the water. If you can see the device without pulling up the rebar check if it is still blinking. If you can’t see well enough to determine this that is ok. If possible, check that thermistors are secure before anticipated high water events.